## Slide 1: Chlorophyll-a calibration curve
Dilutions made from a stock solution measured at 297.7 ug/L
Completed serial dilution from 100 to 5 ug/L in 90% acetone
Vortexed stocks and standards before spiking
Concentration (ug/L) | Counts
100 | 788.625
50 | 389.1875
20 | 199.75
10 | 99.4375
5 | 50.75
0 | 3.0625

## Slide 2: Field validation test
Sampling location close to CORMP EXO2
1.78 ug/L measurement from CORMP
Measurement taken on the surface in an amber bottle
  CORMP measurements deeper
17.89 ug/L average on OF
  Factor of 10 difference, calculation error or sensor issue?

## Slide 3: Potential causes of discrepancy
Calculations
  Spreadsheet for trichromatic chlorophyll method may have calculation mistake
  Problem with calibration curve calculations
  Problem with python script
Environmental factors
  Acetone vs seawater
  Photochemical quenching in seawater
  EXO2 in a slightly different place at a different depth (slightly)
Sensor sensitivity issue

## Slide 4: Solutions
Check all calculations again and have Phil double check them
Calibrate Dr. Loh’s Horiba and test standards on there to determine if it agrees with the spectrophotometer or the OF
Use Van Dorn sampler to sample in the same place and depth as the EXO2
Make chlorophyll standards using CMS highly filtered seawater rather than acetone